Triangle

Environmental, Inc.

### TriTester

### Model: TVLM-001

### Phase II Vapor Recovery

### Air/Liquid Tester

Version 2.96

August 1, 2007

##### OPERATING MANUAL

### TABLE OF CONTENTS

Page

Introduction										1

Safety and Cautions									2

Pretest Checklist									3

Pre and Post Leak Test								4

Performing an A/L Test								5

Maintenance and Care									6

Replacement Parts									6

Technical Support									7

Warranty										7

Warranty Registration Form								8

Appendix:	Air Resources Board June 2, 2004 Approval Letter

Test Report Worksheet

TriTester Parts Diagram

# Introduction

Thank you for purchasing TriTester, Triangle Environmental's Air to Liquid tester Version 2.96.  This state-of-the-art digital tester is used to determine A/L ratios on Phase II vapor recovery systems with unparalleled accuracy.  The system has been certified equivalent for a 2 gallon test "when applied to ANY system for which the 1996 version of ARB TP201.5 is specified or approved."  (June 2, 2004 ARB approval letter)

TriTester has been designed for ease of use, reliability and with the testing technician in mind. Triangle's compliance division tests many thousands of gasoline service stations each year.  We know first hand the challenges facing technicians in today’s ever-changing field of environmental regulations.  We consulted with our technicians to produce a system that meets their needs.  Triangle designed the system with the minimum of embellishments.  TriTester is digitally smart, robust, and compact, crafted into a uniquely advanced design.  Refer to www.trienv.com for additional copies of this manual, ARB approval letter and test report worksheet.

# Overview and Features

- Approved by the Air Resources Board for a 2 gallon test.

- Simple, one button function with audible feedback.

- Automated calculation of results.

- Built in correction factor for the Air Flow Meter.

- Built in calibration expiration date warning.

- The battery is lead-acid rechargeable type with low voltage warning.

- 120 volt ac and 12 volt dc chargers included.

- White backlight LCD display with easy-view for both day and night work.

- Warm-up period is not required.

- Less station downtime, reduced vapor emissions and shorter test times.

Safety and Cautions

TriTester is not currently approved for Class I Division I environments.  Do not use this unit closer than 18 inches to ground level or within 18 inches of a fuel dispenser.  Never disassemble the processor unit or air volume meter for any reason.  Breaking of factory seals or disassembly will void any warranty.

Always follow proper safety procedures when setting-up the unit and when performing tests.  We suggest the following precautions as minimum practice with each use:

- Place warning cones, barricades or caution tape around work area.

- Wear a safety vest, safety glasses and gasoline resistant rubber gloves during testing.

- Ensure that the unit is securely mounted to the test can and that the test can is grounded to the fuel dispenser.

- Always connect the test can grounding strap to the dispenser when conducting tests, and connect the ground strap to the fill manway when draining test can back into underground tank.

- Whenever possible avoid breathing gasoline vapors.  If prolonged exposure is expected a suitable cartridge type facemask is recommended.

- Never use this unit for any purposes other than its intended use.

- Always follow applicable Federal, State, Local and your own company's safety regulations.

CAUTION: Certain nozzle adapters disable the nozzle automatic shut off mechanism, great care should be used when dispensing fuel into the test can. To avoid overfilling and spills, always be aware of the level in the test can when performing tests.  Fire, personal injury and death may occur if these precautions are not followed.

Pretest Checklist

Check the nozzle spout for roundness with the spout ring provided.  If the spout fails the ring test this may invalidate the A/L test.  Nozzle spouts that are damaged such that the V/L adaptor cannot fit over the nozzle spout preclude the use of this test.

Prior to making any adjustments to the vapor recovery system, we recommend conducting a Bulb Test on all the hanging hardware to ensure that the failure is not due to a defective hose, nozzle or breakaway.

The TriTester battery should be recharged when the voltage falls below 11 volts using either the 120 volt or 12 volt chargers supplied with the unit.  The unit is programmed to cease operation if the voltage falls below 10 volts.

Triangle and ARB allow the use of local agency test data result forms provided the form includes the same minimum parameters as ARB’s TriTester form.

WARNING: the unit will not operate after the calibration expiration date. If the calibration date is expired, the display will read "CALIBRATION EXPIRED" and display a phone number for information on the re-calibration procedure.  The unit must be returned to Triangle for factory calibration.

Refer to each particular ARB Executive Order for the System being tested to ensure that the correct A/L nozzle adapter is being used for each type of nozzle tested.  The TriTester nozzle adapter provided with unit can be used with OPW, EZ Flo, EmcoWheaton, and Catlow.  Husky and Healy nozzles require an additional nozzle adapter.

Pre and Post-test Leak Check

The meter must pass a leak test immediately prior to and after conducting A/L verification tests in order for the test to be valid.

The principle of the test is to induce a vacuum on the air flow meter and adapter hose assembly using the hand operated squeeze bulb and observing if the unit holds a vacuum for at least 30 seconds.

Procedure

- Connect the vacuum bulb to meter inlet port using the slip adapter.

- Connect the nozzle adapter hose to air flow meter using the Kamlock fitting.

- Lightly lubricate the ‘O’ rings on the A/L nozzle adapter using grease or other suitable lubricant.

- Slide the blank spout into the appropriate A/L adapter.

- Squeeze the vacuum bulb until the bulb remains compressed, this should take approximately 3 squeezes.

- Once the bulb remains compressed start the stopwatch, if the bulb does not completely re-inflate within 30 seconds, the system has passed the leak test and nozzle testing can begin.

- If the vacuum bulb does not remain compressed, this indicates a leak in the system.  Check all joints and connections for tightness and retest.  If all fittings are tight and the vacuum bulb still does not remain compressed, reverse the vacuum bulb connection to the meter and pressurize the meter with the squeeze bulb.  Use a leak finding solution to detect leaks and repair.  Return to step 5.

## Performing an A/L Test

Turn the on/off switch to the ON position; the following sequence of screens will be displayed:

- Software version number

- Serial number and the calibration expiration date

- Battery voltage status

- Start pumping and press start button when 1 gallon is dispensed

- Press start button at 3 gallons, then stop pumping

- Test results

Connect the A/L adapter to the nozzle spout ensuring a tight connection. Ensure the nozzle spout is in contact with the grounded tank assembly during dispensing.

Start dispensing gasoline, press the "START" button at 1.00 gallon.  You will hear an audible beep each time the button is pressed.

At 3.00 gallons of gasoline, press the button again to stop the test, then release the nozzle trigger to stop pumping.

The display will read gallons per minute, cubic feet recovered, and A/L ratio.

Press button again to start next test.

Note:  The unit does not retain results in memory.  Be sure you have recorded all results before pushing button to go on to the next test.

Results

If the A/L Volumetric Ratio is within the range specified in the applicable CARB executive order, then the test passes.

If the A/L Volumetric Ratio is outside the range specified in the applicable CARB executive order, then the test fails.

#### Note:  If the A/L Volumetric Ratio is outside the range by an A/L value of less than or equal to 0.10, conduct the test two additional times.  Do not make adjustments to the gasoline dispensing equipment, vapor recovery lines or the testing equipment until all three tests have been run.  Calculate the numerical average of the three tests.  If the average A/L value is within the allowable limits, compliance has been verified.  If the average A/L is outside the allowable limits, that fueling point does not comply with the applicable CARB executive order.

#### If the A/L Volumetric Ratio is outside the specified range by a value greater than 0.10, the fueling point does not comply with the applicable CARB executive order and is considered a failure.  No additional tests and averaging are required.

#### Maintenance and Care

Although the unit is sturdily constructed, care should be taken in handling and storage.  The unit should be kept clean, dry and free of any oils or grease.

Check that the inlet port screen on the volume meter is in place and in clean condition.

Examine the nozzle adapter hose for nicks or tears, replace as necessary.

Battery life: The batteries are housed within the unit itself and cannot be removed.  The batteries are rechargeable and the unit comes with a 120 ac and 12 volt dc chargers.

Calibration:  For use in California and in accordance with CARB test procedure TP201.5 (1996 version), the unit must be calibrated annually.  After the expiration date has arrived, the display will read "CALIBRATION EXPIRED."  This is done at Triangle’s facilities in Burbank, California. During the calibration process the unit is checked for accuracy, the battery is renewed and cables are replaced as necessary.  The complete unit with all components must be returned for calibration.

The Air Flow Meter and the Electronic Unit is a calibrated pair and cannot be mismatched with other units.  Disassembly of the unit is not possible without breaking the seals and voiding warranty.

#### TriTester Replacement Parts

#### Description						      	Part Number

Complete System						TVLM-001

Electronic Processor Unit					TVLM-002

Air Flow Meter						TVLM-003

120 volt charger						TVLM-004

Auto charger							TVLM-005

Inlet Port Screen						TVLM-006

Squeeze Bulb							TVLM-007

Squeeze Bulb Slip Adapter					TVLM-008

Spout Ring							TVLM-009

Blank Spout							TVLM-010

Carrying Case							TVLM-011

Spout Adapter							TVLM-012

Technical Support

Triangle can be reached at:				Mailing Address

Phone (800) 767-6004					Triangle Manufacturing, Inc.

Fax     (818) 301-7777					3726 Park Place

Web    http//www.trienv.com				Montrose, CA 91020

# Warranty

All Triangle parts and products are thoroughly inspected and tested at our factory from receipt of raw materials to completion of final assembly.  Triangle Environmental, Inc. (TEI) will, at its option, repair, replace or credit the purchase price of any TEI manufactured product which proves upon examination by TEI to be defective in material and/or workmanship within TWELVE (12) MONTHS from the date of purchase if the original purchaser returns the warranty registration form provided herein. Buyer must return all warranty items to TEI, transportation charges prepaid.

This Warranty does not apply to equipment which has been damaged by misuse, including liquid damage, improper operation or maintenance or which is altered or repaired in any way or for which the factory seals have been broken.

The Warranty provisions contained herein apply ONLY to original purchasers who use the equipment for commercial or industrial purposes.

There are no other warranties of MERCHANTABILITY, FITNESS FOR A PARTICULAR PURPOSE, OR OTHERWISE, and any other such warranties are hereby specifically disclaimed.

TEI assumes NO LIABILITY for labor charges or other costs incurred by Buyer incidental to the service, adjustment, repair, return, removal or replacement of products. TEI ASSUMES NO LIABILITY FOR ANY INCIDENTAL, CONSEQUENTIAL, OR OTHER DAMAGES UNDER ANY WARRANTY, EXPRESS OR IMPLIED, AND ALL SUCH LIABILITY IS HEREBY EXPRESSLY EXCLUDED.

TEI reserves the right to change or improve the design of any TEI equipment without assuming any obligations to modify any equipment previously manufactured.

WARRANTY REGISTRATION

Thank you for purchasing TriTester, the latest in Phase II assisted vapor recovery testing.  We want your unit to be free from defects in material and workmanship, so take a moment to read the warranty, fill out this form and fax or mail to Triangle Environmental, Inc. at:

Triangle Environmental, Inc.

3726 Park Place

Montrose, CA 91020

ATTN: Warranty Registration

Phone (800) 767-6004

Fax (818) 301-7777

Please print

## Name of Company_______________________________________

Street Address__________________________________________

City, State & Zip________________________________________

Phone Number__________________________________________

TriTester Serial #_______________________________________

TriTester Part #_________________________________________

TriTester Version #______________________________________